Instruction Sheet for the Candidate

| Qualification | National Vocational Certificate Level 4 in Mechatronics Technology |
| --- | --- |
| Competency Standard | Calibrate Sensors and Transducers |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name______________________________________________________ Registration/Roll Number_______________________________________ |
| Guidance for Candidate | To meet this standard, you are required to complete the following within 01 Hourtime frame (for practical demonstration & assessment): - Calibrate Capacitive Transducer for Displacement measurement. - Calibrate LVDT Transducer for Displacement measurement. - Calibrate RTD for Temperature measurement. - Calibrate Thermocouple for Temperature Measurement - Calibrate thermistor for temperature measurements. - Calibrate Pressure cell for pressure Measurements. - Calibrate strain gauge for temperature measurement. - Calibrate Stroboscope for Speed Measurement. - Calibrate Stroboscope for Speed Measurement |
| Time: 01 Hour. | During a practical assessment, under observation by an assessor, you are required to: |
| Minimum Evidence Required | - Calibrate Capacitive Transducer for Displacement measurement. - Check connection made to the instrument - Allow the instrument in ON position for 10 minutes for initial warm up - Pull the top plate to Zero position - Adjust the ZERO potentiometer so that the display reads ‘000’ - Move the plate in step of 5 to 10 mm and note down the reading in the tabular column till 50mm. - Calibrate LVDT Transducer for Displacement measurement. - Connect the power supply chord at the rear panel to the 230V 50Hz supply. - Switch on the instrument by pressing down the toggle switch. The display glows to indicate the instrument is ON. - Allow the instrument in ON position for 10 minutes for initial warm-up. - Rotate the micrometer till it reads “20.0”. - Adjust the CAL potentiometer at the front panel so that the display reads “10.0” - Rotate the core of micrometer till the micrometer reads “10.0” and adjust the ZERO potentiometer till the display reads “00.0”. - Rotate back the micrometer core upto 20.0 and adjust once again CAL Potentiometer till the display read (Now the instrument is calibrated for +/- 10.0 mm range. As the core of LVDT moves the display reads the displacement in mm.) - Rotate the core of the micrometer in steps of 1 or 2 mm and tabulate the readings. The micrometer will show the exact displacement given to the LVDT core the display will read the displacement sensed by the LVDT. - Tabulate the readings and Plot the graph Actual V/s indicator readings - Calibrate RTD for Temperature measurement. - Check connection made and Switch ON the instrument by rocker switch at the front panel (The display glows to indicate the instrument is ON.) - Allow the instrument in ON Position for 10 minutes for initial warm-up. - Pore around 3/4th full of water to the kettle and place sensors and thermometer inside the kettle. - Note down the Initial water temperature from the thermometer. - Select the sensor on which the experiment to be conducted through selection switch on the front panel. - Adjust the Initial set Potentiometer in the front panel till the display reads initial water temperature. - Switch on the kept and wait till the water boils note down the reading inn the thermometer and set Final set potentiometer till the display reads boiling water temperature. - Remove the sensor from the boiling water immerse it I the cold water. Set the cold water temperature using initial set potentiometer. - Repeat the process till the display reads exact boiling water and cold water temperature. - Change the water in the kettle with and re heat the water. Now the display starts showing exact temperature raise in the kettle. - Experiment can be repeated for all the three sensors. - Temperature in the thermometer and the indicator readings in steps of 10o C can be tabulated - Calibrate Thermocouple for Temperature Measurement. - Check connection made and Switch ON the instrument by rocker switch at the front panel. - The display glows to indicate the instrument is ON. - Allow the instrument in ON Position for 10 minutes for initial warm-up. - Pore around 3/4th full of water to the kettle and place sensors and thermometer inside the kettle. - Note down the Initial water temperature from the thermometer. - Select the sensor on which the experiment to be conducted through selection switch on the front panel. - Adjust the Initial set Potentiometer in the front panel till the display reads initial water temperature. - Switch on the kept and wait till the water boils note down the reading in the thermometer and set Final set potentiometer till the display reads boiling water temperature. - Remove the sensor from the boiling water immerse it I the cold water. - Set the cold-water temperature using initial set potentiometer. - Repeat the process till the display reads exact boiling water and cold-water temperature. - Change the water in the kettle with and re heat the water. Now the display starts showing exact temperature raise in the kettle. - Experiment can be repeated for all the three sensors. - Temperature in the thermometer and the indicator readings in steps of 100 C can be tabulated. - Calibrate thermistor for temperature measurements. - Check connection made and Switch ON the instrument by rocker switch at the front panel. - The display glows to indicate the instrument is ON. - Allow the instrument in ON Position for 10 minutes for initial warm-up. - Pore around 3/4th full of water to the kettle and place sensors and thermometer inside the kettle. - Note down the Initial water temperature from the thermometer. - Select the sensor on which the experiment to be conducted through selection switch on the front panel. - Adjust the Initial set Potentiometer in the front panel till the display reads initial water temperature. - Switch on the kept and wait till the water boils note down the reading in the thermometer and set Final set potentiometer till the display reads boiling water temperature. - Remove the sensor from the boiling water immerse it I the cold water. Set the cold-water temperature using initial set potentiometer. - Repeat the process till the display reads exact boiling water and cold-water temperature. Change the water in the kettle with and re heat the water. Now the display starts showing exact temperature raise in the kettle. - Experiment can be repeated for all the three sensors. Temperature in the thermometer and the indicator readings in steps of 10o C can be tabulated. - Calibrate Pressure cell for pressure Measurements. - Check connection made and switch ON the instrument by rocker switch at the front panel. The display glows to indicate the instrument is ON. - Allow the instrument in ON Position for 10 minutes for initial warm-up. - Adjust the Potentiometer in the front panel till the display reads “000” - Apply pressure on the sensor using the loading arrangement provided. - The instrument reads the pressure coming on the sensor and display through LED. - Readings the data. - Tabulate the data and % error of the instrument, linearity can be calculated - Calibrate strain gauge for temperature measurement. - Check connection made and Switch ON the instrument by toggle switch at the back of the box. - The display glows to indicate the instrument is ON. - Allow the instrument in ON Position for 10 minutes for initial warm-up. - Adjust the ZERO Potentiometer on the panel till the display roads ‘ OOP’. - Apply load on the sensor using the loading arrangement provided in steps of 100g upto 1 Kg. - The instrument display exact micro-strain strained by the cantilever beam. - Note down the readings in the tabular column. Percentage error in the readings. - Hysteresis and Accuracy of the instrument can be calculated by comparing with the theoretical values. - Calibrate Stroboscope for Speed Measurement. - Check the connections of the equipment. - Select the mode of operation on Xenon flash lamp (Hi/lo). - Make a mark on the rotating wheel. - Now switch on the Xenon flash lamp and adjust that the light exactly projected on the rotating wheel. - Adjust the speed of the motor such that ONE STATIONARY mark is visible on the wheel. Note down the reading in the tabular column. - Now adjust the sped of the motor such that Two Stationary marks are visible and note the readings. - Similarly, for Three and more marks and note the readings. - Calibrate Stroboscope for Speed Measurement. - Connect the turbine flow sensor with indicator marked as flow sensor input. - Connect the two pin of the motor to the instrument. - Vary the flow control potentiometer to any required set level. - Compare the Rotameter reading and digital reading with set reading. - Take reading for different set of flows rate. - Plot the graph of Rotameter Reading with Digital Indicator Reading. |

Self-Assessment Checklist

| Candidate Name |  |
| --- | --- |
| Registration No. |  |
| Qualification | National Vocational Certificate Level 4 in Mechatronics Technology |
| Competency Standard | Calibrate Sensors and Transducers |
| Purpose of Assessment | Formative Assessment |
| Assessment Task | - Identify Various types of Sensors. - Identify Various types of Transducers - Identify Mode of Various Sensors and Transducers. - Operate Oscilloscope for testing sensors and transducers - Test sensor / transducer with Oscilloscope. - Test Temperature Sensor. - Test the IR sensor |

I can……………….

| Performance Criteria | Yes | No |
| --- | --- | --- |
| - Check connection made to the instrument |  |  |
| - Allow the instrument in ON position for 10 minutes for initial warm up |  |  |
| - Pull the top plate to Zero position |  |  |
| - Adjust the ZERO potentiometer so that the display reads ‘000’ |  |  |
| - Move the plate in step of 5 to 10 mm and note down the reading in the tabular column till 50mm. |  |  |
| - Connect the power supply chord at the rear panel to the 230V 50Hz supply. |  |  |
| - Switch on the instrument by pressing down the toggle switch. The display glows to indicate the instrument is ON. |  |  |
| - Allow the instrument in ON position for 10 minutes for initial warm-up. |  |  |
| - Rotate the micrometer till it reads “20.0”. |  |  |
| - Adjust the CAL potentiometer at the front panel so that the display reads “10.0” |  |  |
| - Rotate the core of micrometer till the micrometer reads “10.0” and adjust the ZERO potentiometer till the display reads “00.0”. |  |  |
| - Rotate back the micrometer core upto 20.0 and adjust once again CAL Potentiometer till the display read (Now the instrument is calibrated for +/- 10.0 mm range. As the core of LVDT moves the display reads the displacement in mm.) |  |  |
| - Rotate the core of the micrometer in steps of 1 or 2 mm and tabulate the readings. The micrometer will show the exact displacement given to the LVDT core the display will read the displacement sensed by the LVDT. |  |  |
| - Tabulate the readings and Plot the graph Actual V/s indicator readings |  |  |
| - Connect the power supply chord at the rear panel to the 230V 50Hz supply. |  |  |
| - Switch on the instrument by pressing down the toggle switch. The display glows to indicate the instrument is ON. |  |  |
| - Allow the instrument in ON position for 10 minutes for initial warm-up. |  |  |
| - Rotate the micrometer till it reads “20.0”. |  |  |
| - Adjust the CAL potentiometer at the front panel so that the display reads “10.0” |  |  |
| - Rotate the core of micrometer till the micrometer reads “10.0” and adjust the ZERO potentiometer till the display reads “00.0”. |  |  |
| - Rotate back the micrometer core upto 20.0 and adjust once again CAL Potentiometer till the display read (Now the instrument is calibrated for +/- 10.0 mm range. As the core of LVDT moves the display reads the displacement in mm.) |  |  |
| - Rotate the core of the micrometer in steps of 1 or 2 mm and tabulate the readings. The micrometer will show the exact displacement given to the LVDT core the display will read the displacement sensed by the LVDT. |  |  |
| - Tabulate the readings and Plot the graph Actual V/s indicator readings |  |  |
| - Connect the power supply chord at the rear panel to the 230V 50Hz supply. |  |  |
| - Switch on the instrument by pressing down the toggle switch. The display glows to indicate the instrument is ON. |  |  |
| - Allow the instrument in ON position for 10 minutes for initial warm-up. |  |  |
| - Check connection made and Switch ON the instrument by rocker switch at the front panel. |  |  |
| - The display glows to indicate the instrument is ON. |  |  |
| - Allow the instrument in ON Position for 10 minutes for initial warm-up. |  |  |
| - Pore around 3/4th full of water to the kettle and place sensors and thermometer inside the kettle. |  |  |
| - Note down the Initial water temperature from the thermometer. |  |  |
| - Select the sensor on which the experiment to be conducted through selection switch on the front panel. |  |  |
| - Adjust the Initial set Potentiometer in the front panel till the display reads initial water temperature. |  |  |
| - Switch on the kept and wait till the water boils note down the reading in the thermometer and set Final set potentiometer till the display reads boiling water temperature. |  |  |
| - Remove the sensor from the boiling water immerse it I the cold water. |  |  |
| - Set the cold-water temperature using initial set potentiometer. |  |  |
| - Repeat the process till the display reads exact boiling water and cold-water temperature. |  |  |
| - Change the water in the kettle with and re heat the water. Now the display starts showing exact temperature raise in the kettle. |  |  |
| - Experiment can be repeated for all the three sensors. |  |  |
| - Temperature in the thermometer and the indicator readings in steps of 100 C can be tabulated. |  |  |
| - Check connection made and Switch ON the instrument by rocker switch at the front panel. |  |  |
| - The display glows to indicate the instrument is ON. |  |  |
| - Allow the instrument in ON Position for 10 minutes for initial warm-up. |  |  |
| - Pore around 3/4th full of water to the kettle and place sensors and thermometer inside the kettle. |  |  |
| - Note down the Initial water temperature from the thermometer. |  |  |
| - Select the sensor on which the experiment to be conducted through selection switch on the front panel. |  |  |
| - Adjust the Initial set Potentiometer in the front panel till the display reads initial water temperature. |  |  |
| - Switch on the kept and wait till the water boils note down the reading in the thermometer and set Final set potentiometer till the display reads boiling water temperature. |  |  |
| - Remove the sensor from the boiling water immerse it I the cold water. Set the cold-water temperature using initial set potentiometer. |  |  |
| - Repeat the process till the display reads exact boiling water and cold-water temperature. Change the water in the kettle with and re heat the water. Now the display starts showing exact temperature raise in the kettle. |  |  |
| - Experiment can be repeated for all the three sensors. Temperature in the thermometer and the indicator readings in steps of 10o C can be tabulated |  |  |
| - Check connection made and switch ON the instrument by rocker switch at the front panel. The display glows to indicate the instrument is ON. |  |  |
| - Allow the instrument in ON Position for 10 minutes for initial warm-up. |  |  |
| - Adjust the Potentiometer in the front panel till the display reads “000” |  |  |
| - Apply pressure on the sensor using the loading arrangement provided. |  |  |
| - The instrument reads the pressure coming on the sensor and display through LED. |  |  |
| - Readings the data. |  |  |
| - Tabulate the data and % error of the instrument, linearity can be calculated |  |  |
| - Check connection made and Switch ON the instrument by toggle switch at the back of the box. |  |  |
| - The display glows to indicate the instrument is ON. |  |  |
| - Allow the instrument in ON Position for 10 minutes for initial warm-up. |  |  |
| - Adjust the ZERO Potentiometer on the panel till the display roads ‘ OOP’. |  |  |
| - Apply load on the sensor using the loading arrangement provided in steps of 100g upto 1 Kg. |  |  |
| - The instrument display exact micro-strain strained by the cantilever beam. |  |  |
| - Note down the readings in the tabular column. Percentage error in the readings. |  |  |
| - Hysteresis and Accuracy of the instrument can be calculated by comparing with the theoretical values |  |  |
| - Connect the turbine flow sensor with indicator marked as flow sensor input. |  |  |
| - Connect the two pin of the motor to the instrument. |  |  |
| - Vary the flow control potentiometer to any required set level. |  |  |
| - Compare the Rotameter reading and digital reading with set reading. |  |  |
| - Take reading for different set of flows rate. |  |  |
| - Plot the graph of Rotameter Reading with Digital Indicator Reading. |  |  |

Candidate’s Signature__________________ Assessor’s Signature____________________

Date: _______________________________

Assessors Judgment Guide

| Qualification | National Vocational Certificate Level 4 in Mechatronics Technology |
| --- | --- |
| Competency Standard | Calibrate Sensors and Transducers |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: ______________________Signature: _______________ |
| Assessment Outcome | COMPETENT NOT YETCOMPETENT Name of the Assessor________________________________________________ Assessor’s code: ____________________________________________________ Signature: ________________________________ |

| Assessment Summary (to be filled by the assessor) |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Method |  |  |  |  | Result |  |
| Nature of Activity | Written | Oral | Observation | Portfolio | Role Play | Competent | Not Yet Competent |
| Practical Skill Demonstration |  |  |  |  |  |  |  |
| Knowledge Assessment |  |  |  |  |  |  |  |
| Other Requirement |  |  |  |  |  |  |  |

Observation Checklist

| Assessment Task |  | - Identify Various types of Sensors. - Identify Various types of Transducers - Identify Mode of Various Sensors and Transducers. - Identify operate Oscilloscope for testing sensors and transducers - Identify Test sensor / transducer with Oscilloscope. - Identify Test Temperature Sensor. - Identify Test the IR sensor |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| During the practical assessment, candidate demonstrated the following: |  |  |  | Yes | No | Remarks |
|  | Check connection made to the instrument |  |  |  |  |  |
|  | Allow the instrument in ON position for 10 minutes for initial warm up |  |  |  |  |  |
|  | Pull the top plate to Zero position |  |  |  |  |  |
|  | Adjust the ZERO potentiometer so that the display reads ‘000’ |  |  |  |  |  |
|  | Move the plate in step of 5 to 10 mm and note down the reading in the tabular column till 50mm. |  |  |  |  |  |
|  | Connect the power supply chord at the rear panel to the 230V 50Hz supply. |  |  |  |  |  |
|  | Switch on the instrument by pressing down the toggle switch. The display glows to indicate the instrument is ON. |  |  |  |  |  |
|  | Allow the instrument in ON position for 10 minutes for initial warm-up. |  |  |  |  |  |
|  | Rotate the micrometer till it reads “20.0”. |  |  |  |  |  |
|  | Adjust the CAL potentiometer at the front panel so that the display reads “10.0” |  |  |  |  |  |
|  | Rotate the core of micrometer till the micrometer reads “10.0” and adjust the ZERO potentiometer till the display reads “00.0”. |  |  |  |  |  |
|  | Rotate back the micrometer core upto 20.0 and adjust once again CAL Potentiometer till the display read (Now the instrument is calibrated for +/- 10.0 mm range. As the core of LVDT moves the display reads the displacement in mm.) |  |  |  |  |  |
|  | Rotate the core of the micrometer in steps of 1 or 2 mm and tabulate the readings. The micrometer will show the exact displacement given to the LVDT core the display will read the displacement sensed by the LVDT. |  |  |  |  |  |
|  | Tabulate the readings and Plot the graph Actual V/s indicator readings |  |  |  |  |  |
|  | Connect the power supply chord at the rear panel to the 230V 50Hz supply. |  |  |  |  |  |
|  | Switch on the instrument by pressing down the toggle switch. The display glows to indicate the instrument is ON. |  |  |  |  |  |
|  | Allow the instrument in ON position for 10 minutes for initial warm-up. |  |  |  |  |  |
|  | Rotate the micrometer till it reads “20.0”. |  |  |  |  |  |
|  | Adjust the CAL potentiometer at the front panel so that the display reads “10.0” |  |  |  |  |  |
|  | Rotate the core of micrometer till the micrometer reads “10.0” and adjust the ZERO potentiometer till the display reads “00.0”. |  |  |  |  |  |
|  | Rotate back the micrometer core upto 20.0 and adjust once again CAL Potentiometer till the display read (Now the instrument is calibrated for +/- 10.0 mm range. As the core of LVDT moves the display reads the displacement in mm.) |  |  |  |  |  |
|  | Rotate the core of the micrometer in steps of 1 or 2 mm and tabulate the readings. The micrometer will show the exact displacement given to the LVDT core the display will read the displacement sensed by the LVDT. |  |  |  |  |  |
|  | Tabulate the readings and Plot the graph Actual V/s indicator readings |  |  |  |  |  |
|  | Connect the power supply chord at the rear panel to the 230V 50Hz supply. |  |  |  |  |  |
|  | Switch on the instrument by pressing down the toggle switch. The display glows to indicate the instrument is ON. |  |  |  |  |  |
|  | Allow the instrument in ON position for 10 minutes for initial warm-up. |  |  |  |  |  |
|  | Check connection made and Switch ON the instrument by rocker switch at the front panel. |  |  |  |  |  |
|  | The display glows to indicate the instrument is ON. |  |  |  |  |  |
|  | Allow the instrument in ON Position for 10 minutes for initial warm-up. |  |  |  |  |  |
|  | Pore around 3/4th full of water to the kettle and place sensors and thermometer inside the kettle. |  |  |  |  |  |
|  | Note down the Initial water temperature from the thermometer. |  |  |  |  |  |
|  | Select the sensor on which the experiment to be conducted through selection switch on the front panel. |  |  |  |  |  |
|  | Adjust the Initial set Potentiometer in the front panel till the display reads initial water temperature. |  |  |  |  |  |
|  | Switch on the kept and wait till the water boils note down the reading in the thermometer and set Final set potentiometer till the display reads boiling water temperature. |  |  |  |  |  |
|  | Remove the sensor from the boiling water immerse it I the cold water. |  |  |  |  |  |
|  | Set the cold-water temperature using initial set potentiometer. |  |  |  |  |  |
|  | Repeat the process till the display reads exact boiling water and cold-water temperature. |  |  |  |  |  |
|  | Change the water in the kettle with and re heat the water. Now the display starts showing exact temperature raise in the kettle. |  |  |  |  |  |
|  | Experiment can be repeated for all the three sensors. |  |  |  |  |  |
|  | Temperature in the thermometer and the indicator readings in steps of 100 C can be tabulated. |  |  |  |  |  |
|  | Check connection made and Switch ON the instrument by rocker switch at the front panel. |  |  |  |  |  |
|  | The display glows to indicate the instrument is ON. |  |  |  |  |  |
|  | Allow the instrument in ON Position for 10 minutes for initial warm-up. |  |  |  |  |  |
|  | Pore around 3/4th full of water to the kettle and place sensors and thermometer inside the kettle. |  |  |  |  |  |
|  | Note down the Initial water temperature from the thermometer. |  |  |  |  |  |
|  | Select the sensor on which the experiment to be conducted through selection switch on the front panel. |  |  |  |  |  |
|  | Adjust the Initial set Potentiometer in the front panel till the display reads initial water temperature. |  |  |  |  |  |
|  | Switch on the kept and wait till the water boils note down the reading in the thermometer and set Final set potentiometer till the display reads boiling water temperature. |  |  |  |  |  |
|  | Remove the sensor from the boiling water immerse it I the cold water. Set the cold-water temperature using initial set potentiometer. |  |  |  |  |  |
|  | Repeat the process till the display reads exact boiling water and cold-water temperature. Change the water in the kettle with and re heat the water. Now the display starts showing exact temperature raise in the kettle. |  |  |  |  |  |
|  | Experiment can be repeated for all the three sensors. Temperature in the thermometer and the indicator readings in steps of 10o C can be tabulated |  |  |  |  |  |
|  | Check connection made and switch ON the instrument by rocker switch at the front panel. The display glows to indicate the instrument is ON. |  |  |  |  |  |
|  | Allow the instrument in ON Position for 10 minutes for initial warm-up. |  |  |  |  |  |
|  | Adjust the Potentiometer in the front panel till the display reads “000” |  |  |  |  |  |
|  | Apply pressure on the sensor using the loading arrangement provided. |  |  |  |  |  |
|  | The instrument reads the pressure coming on the sensor and display through LED. |  |  |  |  |  |
|  | Readings the data. |  |  |  |  |  |
|  | Tabulate the data and % error of the instrument, linearity can be calculated |  |  |  |  |  |
|  | Check connection made and Switch ON the instrument by toggle switch at the back of the box. |  |  |  |  |  |
|  | The display glows to indicate the instrument is ON. |  |  |  |  |  |
|  | Allow the instrument in ON Position for 10 minutes for initial warm-up. |  |  |  |  |  |
|  | Adjust the ZERO Potentiometer on the panel till the display roads ‘ OOP’. |  |  |  |  |  |
|  | Apply load on the sensor using the loading arrangement provided in steps of 100g upto 1 Kg. |  |  |  |  |  |
|  | The instrument display exact micro-strain strained by the cantilever beam. |  |  |  |  |  |
|  | Note down the readings in the tabular column. Percentage error in the readings. |  |  |  |  |  |
|  | Hysteresis and Accuracy of the instrument can be calculated by comparing with the theoretical values |  |  |  |  |  |
|  | Connect the turbine flow sensor with indicator marked as flow sensor input. |  |  |  |  |  |
|  | Connect the two pin of the motor to the instrument. |  |  |  |  |  |
|  | Vary the flow control potentiometer to any required set level. |  |  |  |  |  |
|  | Compare the Rotameter reading and digital reading with set reading. |  |  |  |  |  |
|  | Take reading for different set of flows rate. |  |  |  |  |  |
|  | Plot the graph of Rotameter Reading with Digital Indicator Reading. |  |  |  |  |  |
| Competent |  |  | Not Yet Competent |  |  |  |

Knowledge Assessment

| Qualification | National Vocational Certificate Level 4 in Mechatronics Technology |
| --- | --- |
| Competency Standard | Calibrate Sensors and Transducers |
| Purpose of Assessment | Formative Assessment |
| Candidate Details | Name: ______________________________________________________________ Registration/Roll Number: _________________ Candidate Signature: ___________ |
| Assessment Outcome | COMPETENT NOT YETCOMPETENT Name of the Assessor: ________________________________________________ Assessor’s code: ____________________________________________________ Signature of the Assessor: _______________________ |

Candidate’s response is not required to be identical, but similar concepts and/or keywords must be used. Oral questioning may be used to clarify candidate understanding of topic and its application.

| Questions (Candidate confidently answered questions correctly and demonstrated understanding of the topics and their application) |  | Satisfactory | Not Satisfactory |
| --- | --- | --- | --- |
|  | Describe the Calibration Curves |  |  |
|  | Describe validation of results |  |  |
|  | Describe Hysteresis. |  |  |
|  | Describe Linearization |  |  |
|  | Describe Errors in Sensor Measurement. |  |  |
|  | Describe Error due to Improper Zero Reference |  |  |

Feedback to the Candidate

Candidate’s Signature__________________ Assessor’s Signature _________________